## Slide 1
Characterization Methods for Structure-Property Relationships in Clinical Formulations of UHMWPE
Louis Gregory Malito1
1University of California, Berkeley
Adam Kozak2, Stephen Spiegelberg2 PhD, Anuj Bellare3 PhD, Lisa Pruitt1 PhD
2Cambridge Polymer Group, 3Brigham and Women’s Hospital, Harvard Medical School
8th International UHMWPE Meeting
Torino, Italy
Oct. 19th 2017

## Slide 2
Motivation
900,000 TJR annually in U.S. and majority utilize UHMWPE.
Simple mechanical characterization methods are needed for material comparisons and retrieval analysis.
Numerous clinical formulations of UHMWPE in varying crosslink dose, thermal treatment and antioxidant chemistry.
Total Knee Replacement
Total Shoulder Replacement
healthbase.com,, springerimages.com
Total Hip Replacement

## Slide 3
Motivation
What are the Elastic properties of UHMWPE?
What are the post yield properties of UHMWPE?
What are the fracture toughness properties of UHMWPE?
E = 117 to 2020 MPa?
Ultimate Tensile Stress = 33 to 200 MPa
= ( + UTS)/2
(Kurtz et al. Biomaterials 1998)
(Gomoll et al. JORS 2002)
(Bergstrom et al., 2003 Biomaterials)
(Oral et al., 2006 Biomaterials)
(Atwood et al., 2011 JMBBM)
(Bellare et al. JMBBM 2016)
(Ansari et al. JMBBM 2016)
(J-R power law fit)
(Blunting line)
Intersection of J-R curve and
Blunting line is JIC .
(Anderson Fracture Mechanics)

## Slide 4
2 resins (1020/1050)
Range of crosslinking
(Doses: 35-125 kGy)
2 antioxidants:
AO and VE

## Slide 5
Methods: Tensile Testing (Engineering)
ASTM D638 Type IV tensile specimens (n=5) @ 25°C
50 mm/min disp rate
Fracture properties depend on Engineering Ultimate Stress
(Kurtz et al., 1998 Biomaterials)
(Rimnac et al., 1988 PES)
GUR 1020

## Slide 6
Methods: Tensile Testing (True)
ASTM D638 Type IV tensile specimens (n=5) @ 25°C, 50mm/min disp rate.
Dual video extensometer for true stress-strain.
0.2% offset yield for comparison to compression
(Kurtz et al., 1998 Biomaterials)
(Kurtz et al., 2002 Biomaterials)
(Kurtz et al., 2006 Biomaterials)
(Rimnac et al., 1988 PES)
GUR 1020

## Slide 7
Methods: Fracture Toughness
J-R curves using ASTM D6068 and E1820.
1mm/min displacement rate @25°C.
C(T) specimens W=31.75mm, B=15.9mm
EUTS
= ( + UTS)/2
(Paris 1979 ASTM)

## Slide 8
Methods: Microstructure
Crystallinity (XC) through DSC ASTM F2625.
Inter-lamellar spacing (L), lamellar thickness (D), amorphous thickness (A), and specific internal surface (Oac) from from SAXS.
Pearson and Spearman correlation with mechanical properties (median values) to determine relationship between microstructure and bulk properties.
(Atwood et al., 2011 JMBBM)
(Turell & Bellare 2004 Biomaterials)

## Slide 9
Results: Tensile Stress Strain
Linear regression from 0.0005 to 0.009 true strain produces largest difference between moduli of UHMWPE material formulations.
Highest R2 values produced this way (0.95-0.99).
Relative STD is not increased using this method.
(Oral et al., 2006 Biomaterials)
(Atwood et al., 2011 JMBBM)
Method can elucidate or hide material properties!!!

## Slide 10
True Ultimate Tensile Stress decreases with cross-linking dosage in each material group.
True Ultimate Tensile Strain decreases with cross-linking dosage in each material group.
Energetic Toughness (ET) decreases with cross-linking dosage in each material group.

## Slide 11
Results: Fracture Toughness
E = 865.4 MPa
14.1 MPa
EUTS = 49.4 MPa
E = 865.4 MPa
14.1 MPa
TUTS = 158.8 MPa
= 86.5 MPa
(Rimnac et al., 1988 PES)
(Pascaud et al., 1997 PES)
(ASTM E1820)

## Slide 12
Results: Fracture Toughness
Tearing Modulus (T) (Flow=Metal), no real change between blends
KJIC (Flow=Metal) ranges between 4.4 to 5.7 MPa√m
Varadarajan and Rimnac found differences of -45% to -30% between JminQ from CTOD versus JQ from blunting line for cross-linked UHMWPE
(Varadarajan & Rimnac 2008 Polymer)
It’s all about the METHOD!

## Slide 13
Results: Microstructure & Correlations
& TUT Stress |  = 0.628
& TUT Strain |  = 0.8
ET |  = 0.78
A & TUT Strain |  = 0.601
L & EUT Stress |  = 0.67
D & EUT Stress |  = 0.64
 |  = 0.628
 |  = 0.8
 |  = 0.78
A & TUT Strain |  = 0.601
L & EUT Stress |  = 0.67
D & EUT Stress |  = 0.64
Spearman Rank Correlation
from 45.6 – 52.9% across material formulations
D from 21.3 – 25.9nm across material formulations
(Atwood et al., 2011 JMBBM)
& KJIC (Flow=Metal) |  = 0.912
& dJ/dΔa |  = 0.946
EY Stress & KJIC (Flow=Metal) |  = 0.943
 |  = 0.912
 |  = 0.946
EY Stress & KJIC (Flow=Metal) |  = 0.943
Pearson Correlation Coefficient
Spearman Rank Correlation
p ≤ 0.05 (all correlations using median values)

## Slide 14
Mechanical testing of UHMWPE needs to be standardized as method can elucidate or hide material properties.
Linear regression from 0.0005 to 0.009 true axial strain, from true stress-strain data offers the best method of analyzing elastic modulus.
Multi-specimen J-R data passes validity criteria.
Using EUTS as can over exaggerate JIC and KJIC , more conservative to use True Ultimate Tensile (TUT) Stress with ASTM E1820 flow stress. Using true data with E1820 approach comparable to previous findings with CTOD.
correlate with TUT Stress, TUT strain, ET, and dJ/dΔa.
D and EY Stress was found to correlate to KJIC (Flow=Metal) implying that altering the crystalline phase of UHMWPE can increase fracture toughness.
Conclusions
(Varadarajan & Rimnac 2008 Polymer)

## Slide 15
We would like to thank Orthoplastics, Quadrant and DePuy for supplying materials
Funding for this research was provided by the Lawrence Talbot Professorship endowment, and the Ian Finnie graduate mechanical behavior of engineering materials fellowship.
Acknowledgements